#L27

Q: How do I use external clock?

#### External Clock

Use of the External Clock feature in ETSNT permits users to synchronize the ETS system clock with certain other instruments. This feature can be employed in two ways: 1) ETS clock drives a clock input line to another instrument; or, 2) A clock is provided to the ETS system to be used as the Test Rate clock.

##### ETS Provides the Clock

This application is simple. Just connect a standard BNC cable to the Clock Out BNC connector on the front of the ETS system. Connect the cable to the 50 ohm input of a scope; you should see an "active lo" TTL clock. Clock pulses should be "lo" for 10ns, with the leading edge of each pulse providing a rate of the same frequency as the Test Rate entered on your Run Setup window. Here's an example:

Accessing "Clock Out" for External Use

ETS Uses Your Clock

If your goal is to provide and control a clock source that will be used as the "Test Rate" in Run Setup, use the Clock In BNC connector on the front of the ETS system. Your clock source needs to provide a TTL active-lo or "lo-going" pulse of 10ns width, regardless of the input frequency you are using. Use the Test ExtClock button to see the pulse rate as sampled by the tester.

Running ETS on an External Clock

ETS Models 270, 370 and 570

For these models, the external clock input must be twice the rate of the desired operating clock rate (i.e., if external clock rate is 20MHz, then the actual clock rate used by the system is 10MHz). This is not a bug, but is by design. The divided frequency is displayed when the Test ExtClock button is clicked, and is also                                              the time base for calculating the delays as a percentage. The ETS 770 does not divide the clock, but uses the actual sampled clock.